ContinuousAO User’s Guide

Current version of document : 1.0

Last  modification  le 30/07/2026

History of modifications

| Date | Revision | Description | Author |
| --- | --- | --- | --- |
| 17/11/04 | 1.0 | Initial Version | G.Abeillé |

Table of contents

1	Applicative goals and main functionnalities of the DeviceServer	2

1.1	Introduction	2

2	Technical scheme of the application	2

2.1	Board structure	2

2.2	Patch panel	3

3	TANGO software interfaces	3

3.1	Properties	3

3.2	Attributes	5

3.3	Commands	6

3.4	Exceptions	6

4	Performances	6

5	Software Setup	6

5.1	Environment requisites	6

5.2	Installation of the boards	7

5.3	Getting the board number	7

5.4	Tango DeviceServer installation	7

# Applicative goals and main functionnalities of the DeviceServer

## Introduction

This device, ContinuousAO, performs continuous analog output (waveform generation), simultaneously on several channels (from 1 to 8 channels). It is used with the cPCI board SAO_2502 from ADLink. The waveform generation can be started and stopped by software or by external signals. These signals are called triggers. With this device, we can use a start trigger and a stop trigger.

# Technical scheme of the application

## Board structure

This board contains 8 output channels. These channels are distributed in two groups. Channels 0, 1, 2, 3 belong to group A and channels 4, 5, 6, 7 belong to group B. Most of the configuration is common for both groups. Some parameters are different for the 2 groups. Here are some requisites to know before using this device.

### Memory organisation

### There are two onboard memories named FIFO (First In, First Out), one for group A and one for group B. The output samples are transferred to this memory, and then outputted at a precise rate, totally independently from the CPU loading. But, these FIFO have a limited size (~8Ksamples per group), and if a period of the output signals does not fit in the FIFO, the data must be transferred repeatedly between the memory of the CPU and the FIFO with the DMA (Direct Access Memory) protocol. In this case, if the CPU is overloaded, it can happen that data has not been transferred in time for outputting, and old data, previously saved in the FIFO, will be outputted. This phenomenon is named underrun.

### It is important to note that if the data cannot be loaded onto the FIFO, then the number of channels activated on each group must be the same. Otherwise, an exception will be thrown.

### Configuration

If the external digital trigger is used (for start trigger), a trigger signal must be connected for each group, A and B (see patch panel).

The data can be outputted in a certain range that is configurable through the properties of the device. There are two cases:

- The board is configured to use an internal reference of 10 volts. Then, the output range can be configured to be bipolar to output data between -10 and +10 volts. Or it can be configured to be unipolar to output data between 0 and 10 volts.

- The board is configured to use external references. These external references are different for group A and B. Group A uses the reference AORefA and group B uses AORefB. The signals connected to these pins must be between -10 and +10 volts. The range can be also bipolar or unipolar. So, in the unipolar case, the output range for group A’s channels is 0 to +AORefA volts and 0 to +AORefB for group B’s channels. If the range is bipolar, it is – AORefA to + AORefA for group A and – AORefB to +AORefB for group B.

This range configuration is performed through the following properties of the device (see next chapter):

- AORefSource: source of the reference, internal or external.

- AORefA and AORefB: voltage references if external reference source selected.

- Polarity: bipolar or unipolar.

## Patch panel

In progress.

# TANGO software interfaces

## Properties

The properties contain the fixed parameters of the DeviceServer. They are filled by the user at install time, and are not supposed to be modified during runtime.

POGO generated documentation

| boardNum | The board num in the cPCI chassis. |
| --- | --- |
| Frequency | The output frequency, at which each sample (for one channel) is outputted. (in Hz) |
| StartTrigger | Put 1 to enable the start trigger, 0 otherwise. See StartTriggerSource to configure it. |
| StopTrigger | Put 1 to enable the stop trigger, 0 otherwise. See StopTriggerSource to configure it. |
| StartTriggerSource | Configure the type of start trigger. Can be digital or analog. The possible values are: - DTRIG: Digital trigger. See DTRIGPolarity to configure it. - ATRIG: Analog trigger. See ATRIGSelection and ATRIGLevel. NB: Only one trigger (start or stop) can use the analog trigger. |
| StopTriggerSource | Configure the stop trigger. Can be digital or analog. The possible values are: - ATRIG: Analog trigger. See ATRIGSelection and ATRIGLevel - AFI0: Digital line AFI0. - AFI1: Digital line AFI1. NB: Only one trigger (start or stop) can use the analog trigger. |
| ATRIGSelection | If the analog trigger is used (StartTriggerSource or StopTriggerSource configured with ATRIG), configure the type of selection. A trigger can occur when the analog signal passes below a certain level or above a certain level. The possible values are: - BELOW: trigger occurs when passing below a level. - ABOVE: trigger occurs when passing above a level. See ATRIGLevel to parameter the level. |
| ATRIGLevel | The level, in volts (-10 to +10), at which the analog trigger will occur. (8 bits resolution. |
| DTRIGPolarity | The digital trigger polarity. Can be RISING, to detect rising edges, or FALLING, to detect falling edges. |
| channel0Enable channel1Enable channel2Enable channel3Enable channel4Enable channel5Enable channel6Enable channel7Enable | Properties to enable the desired channels. Put 1 in the property to enable the channel, 0 otherwise. (E.g. If channel0Enable if 1, data will be outputted on channel 0). |
| AORefA | The value of the reference connected to the pin AORefA. It used to configure the output range of channels 0, 1, 2, 3. The possible values are -10.0 to 10,0 volts. Only used when the property AORefSource is EXTERN. |
| AORefB | The value of the reference connected to the pin AORefB. It used to configure the output range of channels 4, 5, 6, 7. The possible values are -10.0 to 10,0 volts. Only used when the property AORefSource is EXTERN. |
| Polarity | To configure the output range. Can be: - BP for bipolar (output between –ref to +ref volts) - UP for unipolar(output between 0 to +ref volts). |
| AORefSource | To configure the reference used for the output range. - EXTERN: The reference used is the external pins AORefA and AORefB. - INTERN: The reference is internal. Its value is 10 volts. |
| BufferDepth | The buffer depth for one output channel, i.e. the number of samples for one period of the signal and for one channel. |
| Channel0Waveform Channel1Waveform Channel2Waveform Channel3Waveform Channel4Waveform Channel5Waveform Channel6Waveform Channel7Waveform | Properties used for persistency of attributes (cf. attributes §). |

If the properties are not present in database, the following default values are loaded:

- BoardNum =  0;

- Frequency = 50000.0; //50 kHz

- StartTrigger = 0; //no start trigger

- StopTrigger = 0;  //no stop trigger

- StartTriggerSource = DTRIG;

- StopTriggerSource = AFI1;

- ATRIGSelection = BELOW;

- ATRIGLevel = 5.0;

- DTRIGPolarity = RISING;

- AORefA = 10.0;

- AORefB = 10.0;

- Polarity = BP;

- AORefSource = INTERN;

- Channel0Enable = 1;

- Channel1Enable = 0;

- Channel2Enable = 0;

- Channel3Enable = 0;

- Channel4Enable = 0;

- Channel5Enable = 0;

- Channel6Enable = 0;

- Channel7Enable = 0;

- BufferDepth = 4096;

## Attributes

POGO generated documentation

- channel0: spectrum that contains one period of the current outputted signal on output channel 0.

- …

- channel7: spectrum that contains one period of the current outputted signal on output channel 7.

- errorCounter: Since the waveform generation is performed in a separate thread, some errors can occur and will not be caught in the device. So errorCounter increments each time an error occurs.

- useBoardFifo: Inform the user if the output data has been loaded in the board’s FIFO. It depends of total number of samples needed. If data is loaded on FIFO, userBoardFifo is equal to 1, 0 otherwise. If the data could not be loaded in the FIFO, there is a risk of underrun is the CPU is overloaded.

The attributes channel0, channel1, channel2, channel3, channel4, channel5, channel6, channel7 are persistent. Each time they change, there contains are saved in database (in properties channel0Waveform…channel7Waveform). When the device restarts, the values saved in database are loaded. But if the size of the spectrum saved in the database does not have the same size as specified in property BufferDepth, a default waveform will be loaded. A default waveform is also loaded if the waveform is not set in database.

## Commands

POGO generated documentation

| Start | Start the waveform generation on the activated channels. (If a start trigger is activated, the board will wait for the trigger to start the generation). |
| --- | --- |
| Stop | Stop the current started generation. |
| SetAOScaledData | Command to give one period of the output signal for a specified channel. The samples are given in volts. The voltages must be coherent with the ranges, otherwise the signals will not be correct (e.g. do not give a signal between -10 and +10 if the configured range is 0 to 5 volts). This waveform will be saved in database. When launching the device, if waveforms are not in database or if no data has been given through this command, default waveforms are generated (sinusoids on group A and square signals on group B). NB: When calling this command, the current running generation is stopped. The user must call Start command to start the generation with new waveforms. |

NB: The generation must always be started by user (i.e. it is never done in command Init or SetAOScaledData), because generating a wrong signal on a specific hardware can be dangerous.

## Exceptions

### If the number of channels configured is not the same in group A and group B, and that data does not fit in the FIFO, an exception occurs. This is a hardware limitation, to perform DMA transfer, the number of channels activated in group A and in group B must be the same.

If a buffer is given to the device through the command SetAOScaledData with a different size from the bufferDepth gave in the properties, an exception will arise.

If the board is not well installed (or not present in the chassis) or if the property boardNum is not configured with the right value (see below for more details for retrieving the board number), the device cannot start.

# Performances

When data fits in the onboard FIFO, it is more efficient than DMA transfer that can cause underrun. So, if possible, the user must care about fitting data into the FIFO (the user is informed with attribute useBoardFifo).

# Software Setup

## Environment requisites

This device works under Windows, it not yet tested under Linux.

The device needs at least the version 4 of TANGO.

## Installation of the boards

Two drivers, D2K-DASK and PCIS-DASK for the ADLink boards are needed. These drivers exist under Windows and Linux.

- Installation under windows:

  - Install the drivers.

  - Shut down the chassis.

  - Insert the board in a slot of the chassis.

  - Restart the chassis.

  - While restarting windows, the device should be detected and install automatically by the OS.

  - If the board is correctly installed, it should appear in the Device Manager.

## Getting the board number

The board number (cf. properties) depends on different parameters. It depends on the number of boards of the same type in the chassis. And it also depends on the slot where the required board is.

If there is only one board of the same type in the chassis, the board number is always 0. If there are ‘n’ boards of the same type, they will be numbered in function of there position. The board located in the slot with lower number (generally the closest slot to the CPU) will get the board number 0. The second board will get the value 1, the third board will get 2, etc…

Be careful, that if boards are taken away or added to a CompactPCI crate, the numbering scheme will change.

## Tango DeviceServer installation

The device is not dependant of other devices. It can be started and displayed with the common procedure for a device server. To use the required board, it is mandatory to fill the property “BoardNum””. Then the device can start with the other properties set to default values.