DeviceServer AcquireWaveformLecroy

Author : X.Elattaoui

Current version : 1.0

Creation Date:  September 2004

Last modification

Historical modifications

| Date | Current Revue | Description | Author | Approver |
| --- | --- | --- | --- | --- |
| 09/2004 | 1.0 | Initial version | X.Elattaoui | A. Buteau |

1	Goal	2

2	Hardware and Software Structure	2

2.1	Hardware architecture	2

2.2	Software architecture of the DeviceServer	2

2.3	Internal Classes description :	2

2.4	Usage principle	2

3	Known Limitations of the DeviceServer AcquireWaveformLecroy	3

3.1	Possible enhancements	3

3.2	Incompatibility with RemoteDSO application	3

4	Tango Software Interface	3

4.1	Properties	3

4.2	The attributes	3

4.3	The commands	4

5	How to use the AcquireWaveformLecroy DeviceServer	4

5.1	First use of the AcquireWaveformLecroy DeviceServer	4

# Goal

The main purpose of this DeviceServer is to transfert  waveforms acquired by Lecroy scopes in the Tango world.

Both raw and scaled data of a waveform acquired on any channel of a Lecroy scope are managed by this DeviceServer.

Acquisitions has to be parametrized on the scope itself, as this deviceServer doesn’t permit to change any acquisition nor scope parameters.

# Hardware and Software Structure

## Hardware architecture

One or several Lecroy scopes as WavePro or WaveRunner series, with an Ethernet interface.

## Software architecture of the DeviceServer

The AcquireWaveformLecroy device server uses internally  the following classes

## Internal Classes description :

- SocketLecroy class:

There is only one instance of this class for 1 DeviceServer .

This class manages the socket communication  (writes/reads data ) with  the Lecroy scope.

- Waveform class :

This class sends the commands to acquire the waveform data, extract significant fields from raw data sent by the scope.

## Usage principle

- The user gives to the device :

  - The channel name : is composed by a letter followed by a number

Ex :« C2 for » channel 2, « F1 » for Fonction F1 and so on.

- The IP address of the device

- The user gets the 2 waveforms :

  - The raw data

  - And the vertically scaled data

# Known Limitations of the DeviceServer AcquireWaveformLecroy

## Possible enhancements

The horizontal scaled data are not available yet (but can be developped if required)

More information describing the acquisition (ex ….)  may be extracted if required.

## Incompatibility with RemoteDSO application

The scope can communicate remotely through a TCP/IP link with only 1 client application.

Pratically this means that you cannot use simultaneously :

- The RemoteDSO application provided by Lecroy for a remote control of their scopes

- A TANGO DeviceServer

Anyway RemoteDSO is not a very convenient application and using VNC to remote access the scope is a much convenient solution.

It also means, that you cannot start 2 Tango DeviceServers communicating with the same scope (it would be a strange idea to do so, but users of controls systems are sometimes strange people)

# Tango Software Interface

## Properties

The properties contain non removable parameters of the device. The user sets this parameter when he starts the DServer, it can’t be changed during the process.

| IPaddress | The IP address of the Lecroy scope |
| --- | --- |
| ChannelName | The name of the channel on which the waveform data will be acquired. Ex : “C2” , “F1” … |

## The attributes

| rawWaveformData | Represents the raw data of the acquired waveform |
| --- | --- |
| verticalScaledData | The vertical scaled data calculated as follow : (VERTICAL_GAIN * rawdata) – VERTICAL_OFFSET |
| waveArray1 | The length of the data array 1 |
| waveArrayCount | The total length of the acquired data. (length of data array 1 and data array 2) |
| nominalBits | Is the ADC resolution. For simple waveform its 8 bits but for complex waveform (as FFT …) its 10 up to 12 bits |
| horizontalInterval | Sampling interval for time domain waveforms |
| horizontalOffset | It’s the “delay” |
| verticalGain | The vertical gain |
| verticalOffset | The vertical offset |
| triggerTime | The time of the acquisition in format MM/DD/YEAR |

## The commands

No commands

# How to use the AcquireWaveformLecroy DeviceServer

## First use of the AcquireWaveformLecroy DeviceServer

### Declare the AcquireWaveformLecroy DeviceServer on the Tango software bus

You must declare 1 Dserver per scope :Ex: “ds_ AcquireWaveformLecroy/7100”

### Declare as many Device as channel to acquire

You must declare 1 Device per channel :

Ex: “ANS-C08/DG/SCOPE7100-F1” for the channel F1

“ANS-C08/DG/SCOPE7100-C3” for the channel 3

- Give them the following  mandatory properties :

  - IPaddress : the IP address of the scope

  - ChannelName : the name of the channel of the waveform to acquire