| Diffusion: |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Distribution list Copy to: |  |  |  |  |  |  |  |  |
| Date | Writer | Reviewer | Approver |  | Modifications |  |  | Version |
| 14/10/2016 | Falilou THIAM | Sonia MINOLLI |  |  | Initial release |  |  | 1 |

SGonPPMAC

User Manual

Table of contents

1	Introduction	4

2	Controller configuration	5

2.1	Reminders – SGonPPMACBox interface	5

2.1.1	Device description	5

2.1.2	Device properties	5

2.1.3	Device static attributes	5

2.1.4	Device commands	6

2.1.5	Device states	6

2.2	Reminders – SGonPPMACAxis interface	7

2.3	Reminders – SGonPPMACComposedAxis interface	7

2.3.1	Device description	7

2.3.2	Device Properties	7

2.3.3	Device static attributes	7

2.3.4	Device commands	8

2.3.5	Device states	8

2.4	Operating concept	9

2.5	Axes configuration / Initialization	10

2.6	Motion Controller Management	11

2.7	ComposedAxis Management	13

3	ANNEXES	16

3.1	Details of box device status	16

3.2	Details of axis device status	16

Illustrations table

Figure 1 : SGonPPMACBox in Operator View	12

Figure 2 : SGonPPMACBox in Expert View	12

Figure 3 : SGonPPMACComposedAxis in Operator View	14

Figure 4 : SGonPPMACComposedAxis added attributes in Expert View	15

Tables table

Aucune entrée de table d'illustration n'a été trouvée.

# Introduction

The purpose of this document is to provide additional information on the usage of the SGonPPMAC motion controller Devices SGonPPMACBox, SGonPPMACComposedAxis.

The base of the Device documentation remains the html files, so the user should first look at the html documentation.

# Controller configuration

## Reminders – SGonPPMACBox interface

Here is the description of the SGonPPMACBox Tango interface from its html documentation: properties, attributes, commands and states.

### Device description

Class to manage the Power PMAC motion controller through the PowerPMACLib library.

The main functions are:

- Manages connection to the motion controller,

- Gets controller general parameters,

- Provides general commands,

- Manage the number of deployed instances of composed axes.

### Device properties

| Property name | Property type | Description | Possible values | Default value |
| --- | --- | --- | --- | --- |
| IPAddress | Tango::DEV_STRING | IP address of the underlying power brick. This property has no default value and must be specified. Ex: 192.168.12.3 |  | No default value |
| TCPPort | Tango::DEV_ULONG | Power brick listening TCP port. |  | 22 |
| StdDataPollingPeriod | Tango::DEV_DOUBLE | Polling period for standard data reading, in ms. If set to null value, the standard data reading is synchronous. For expert only. |  | 0 |
| CriticalDataPollingPeriod | Tango::DEV_DOUBLE | Polling period for critical data reading (axis state and position), in ms. If set to null value, the critical data reading is synchronous. For expert only |  | 100 |
| NbOfComposedAxes | Tango::DEV_USHORT | Number of deployed instances of composed axes | From 0 to ‘n’ Where ‘n’ is the number of deployed instances of composed axes. | 1 |

### Device static attributes

| Attribute name | Data Type | R/W Type | Expert |
| --- | --- | --- | --- |
| cpuTemperature: Current CPU temperature in degrees. | DEV_DOUBLE | READ | No |
| runningTime: Time from last power-on to present, in seconds. | DEV_DOUBLE | READ | No |
| phaseTaskUsage: Percentage of averaged phase tasks calculation by task time. | DEV_DOUBLE | READ | Yes |
| servoTaskUsage: Percentage of averaged servo tasks calculation by task time. | DEV_DOUBLE | READ | Yes |
| rtIntTaskUsage: Percentage of averaged real-time interrupt tasks calculation by task time. | DEV_DOUBLE | READ | Yes |
| bgTaskUsage: Averaged calculation time for background task scans. | DEV_DOUBLE | READ | Yes |

### Device commands

| Command name | Description | Level | Argument In | Argument Out |
| --- | --- | --- | --- | --- |
| Init | This commands re-initialise a device keeping the same network connection. | Operator | DEV_VOID | DEV_VOID |
| State | This command gets the device state (stored in its device_state data member) and returns it to the caller. | Operator | DEV_VOID | DEV_STATE |
| Status | This command gets the device status (stored in its device_status data member) and returns it to the caller. | Operator | DEV_VOID | CONST_DEV_STRING |
| Abort | Stops all movements and programs | Operator | DEV_VOID | DEV_VOID |
| Reset | Resets the Power Brick | Expert | DEV_VOID | DEV_VOID |
| Reboot | Reboots the Power Brick operating system | Expert | DEV_VOID | DEV_VOID |
| ExecLowLevelCmd | Execute low level command | Expert | DEV_STRING | DEV_STRING |
| GetFirmwareVersion | Returns current firmware version | Expert | DEV_VOID | DEV_STRING |
| GetUCodeVersion | Returns current microcode version | Expert | DEV_VOID | DEV_STRING |

### Device states

CLOSE: communication closed with controller.

OPEN: communication opened with controller.

ALARM: minor error occurred.

FAULT: fatal error occurred.

INIT: Initialization in progress.

## Reminders – SGonPPMACAxis interface

This interface is the same as PowerPMACAxis, which was made for the PowerPmac purpose (DeltaTAU controller). SGonPPMACAxis is in our devices architecture because SGonPPMAC has been coded based on PowerPMac devloppement. In SGonPPMAC case, SGonPPMACAxis is not used (there is no single axis available).

## Reminders – SGonPPMACComposedAxis interface

Here is the description of the SGonPPMACComposedAxis Tango interface from its html documentation: properties, attributes, commands and states.

### Device description

Class to manage each composed axis defined in controller.

A composed axis is set of virtual axes, representing coupled motors.

The main functions are:

- Manages underlying motor(s),

- Provides interface to set/get configuration parameters for composed axis,

- Provides interface to set/get axis state.

### Device Properties

| Property name | Property type | Description | Possible values | Default value |
| --- | --- | --- | --- | --- |
| AxisDefinition | Array of string | For composed axis, list of: <Coordinate System>::<axis name>::<user unit> Ex: 50::X::slitGap::mm Linear possible units: m, mm, µm, nm Angular possible units: rad, mrad, deg, mdeg |  | No default value |
| MotionPrograms | Array of string | List of available motion programs for this CS. Format: <motion program id>::<description> Ex: 1::trajectory test 2::wash machine |  | No default value |
| DisabledCmds | Array of string | List of disabled commands | on off mpstart disable | No default value |

### Device static attributes

| Attribute name | Data Type | R/W Type | Expert |
| --- | --- | --- | --- |
| velocity: Velocity of linear movement, in user unit/s | DEV_DOUBLE | READ_WRITE | No |
| accelerationTime: Current CS acceleration, in second | DEV_DOUBLE | READ_WRITE | Yes |
| decelerationTime: Current CS deceleration, in second | DEV_DOUBLE | READ_WRITE | Yes |
| accelSCurve: Current S-curve portion of the acceleration, in second | DEV_DOUBLE | READ_WRITE | Yes |
| <label>: Current axis position for the <label> axis, in user unit | DEV_DOUBLE | READ_WRITE | No |
| <label>offset: Stub offset on axis position for the <label> axis, in user unit | DEV_DOUBLE | READ_WRITE | Yes |
| <label>CompensationEnabled: Allows user to enable/disable position <label> compensation table | DEV_BOOLEAN | READ_WRITE | Yes |
| movingMode: Type of movement, among: 0 = rapid, 1 = linear, 2 = spline, 3 = PVT, 4 = circle. | DEV_USHORT | READ_WRITE | Yes |
| freeze: If set to true, position setpoints are memorized and applied on true -> false transition. If false, a position setpoint is immediately applied. | DEV_BOOLEAN | READ_WRITE | No |
| motionProgramState: Current motion program state | DEV_STRING | READ | No |

### Device commands

| Command name | Description | Level | Argument In | Argument Out |
| --- | --- | --- | --- | --- |
| Init | This commands re-initialise a device keeping the same network connection. | Operator | DEV_VOID | DEV_VOID |
| State | This command gets the device state (stored in its device_state data member) and returns it to the caller. | Operator | DEV_VOID | DEV_STATE |
| Status | This command gets the device status (stored in its device_status data member) and returns it to the caller. | Operator | DEV_VOID | CONST_DEV_STRING |
| On | Power on all the underlying motors. | Operator | DEV_VOID | DEV_VOID |
| Off | Power off all the underlying motors. | Operator | DEV_VOID | DEV_VOID |
| Stop | Stop current axis(axes) movement. | Operator | DEV_VOID | DEV_VOID |
| Off_omega | Power off the OMEGA underlying motors | Expert | DEV_VOID | DEV_VOID |
| Off_others | Power off the underlying motors except OMEGA axis | Expert | DEV_VOID | DEV_VOID |
| Disable (Stays in device interface but not used anymore, see property DisabledCmds) | Disable composed axis Device and enable associated simple axis Devices. | Expert | DEV_VOID | DEV_VOID |
| InitializeReferencePosition | User must deactivate all compensation tables before to hit initializeReferencePosition otherwise, an error will occur. Start position initialization process. | Operator | DEV_VOID | DEV_VOID |
| MotionProgramStart | Start the specified motion program id. | Operator | DEV_LONG | DEV_VOID |
| MotionProgramPause | Pause the current motion program. | Operator | DEV_VOID | DEV_VOID |
| MotionProgramResume | Resume the previously paused motion program. | Operator | DEV_VOID | DEV_VOID |

### Device states

OFF: underlying motor is OFF.

STANDBY: underlying motor is ON, no motion in progress.

MOVING: underlying motor is moving.

ALARM: minor error occurred.

FAULT: fatal error occurred.

DISABLE: Axis not controlled by device.

CLOSE: Referencing not done or in error.

## Operating concept

First of all, the user has to define the 6 axes of the SmarGon ComposedAxis. Those 6 axes can’t be used separately.

The user should define the number of Composed Axis, in our situation there is only one.

- ComposedAxis: Represents the set of coupled Axes for the SGonPPMac.

|  | Beware: The virtual axis names should be different. |
| --- | --- |

The SGonPPMAC system is composed by:

- A unique TANGO Device to manage the SGON PPMAC motion controller: SGonPPMACBox

  - Manages connections to the motion controller,

  - Gets controller general parameters (temperature, CPU data, firmware versions,…),

  - Provides general commands (motion abort, reset, os reboot, low level commands, …),

  - Contains the property for the number of deployed instances of composed axes.

|  | Beware: For SGonPPMACBox Device, the NbOfComposedAxes property is mandatory: it should be set 1. |
| --- | --- |

- Old TANGO Device to manage each simple axis defined in controller: SGonPPMACAxis

  - In our case, SGonPPMMACAxis can’t work. SmarAct provides a specific API, the access for only one Axis is not available in this version.

  - This device class is still automatically created by SGonPPMACBox but not instantiated.

- A unique TANGO Device to manage SmarGon system: SGonPPMACComposedAxis

  - Manages underlying motor(s): ON/OFF/STOP…

  - Provides interface to set/get configuration parameters for composed axis (velocity, acceleration, deceleration)

  - Provides interface to set/get axis state (state, position),

  - Provides interface to manage CS motion program (note that only one motion program at a time can be running for a specified CS).

## Axes configuration / Initialization

In the SGonPPMACBox Device properties, the user describes the connection configuration, for example:

| Property name | Property content |
| --- | --- |
| IPAdress | 192.168.0.1 |
| TCPPort | 22 |
| NbOfComposedAxes | 1 |
| StdDataPollingPeriod | 0 |
| CriticalDataPollingPeriod | 100 |

|  | Beware: The NbOfComposedAxes property is mandatory: it has to be set to 1. |
| --- | --- |

For SGonPPMACComposedAxis Device, the user has to describe the axis configuration in Device properties, which is:

| Property name | Property content |
| --- | --- |
| AxisDefinition | 1::A::omega::deg 1::B::phi:: deg 1::C::chi::deg 1::X::x::mm 1::Y::y::mm 1::Z::z::mm |

|  | Beware: The rotation on axis OMEGA is inverted compared to the other axes. |
| --- | --- |

## Motion Controller Management

The available operations for SGonPPMACBox Device are:

- Read current motion controller state,

- Send low level commands (reset, abort, reboot, …).

|  | Beware: If Box Device is in FAULT state, the composed axis will be in FAULT state. |
| --- | --- |

Acquisition parameters:

- cpuTemperature  (Operator)

- runningTime   (Operator)

- phaseTaskUsage  (Expert)

- servoTaskUsage  (Expert)

- rtIntTaskUsage    (Expert)

- bgTaskUsage       (Expert)

- controller status (cf. § 3.1).

Available Commands:

- Abort, Init, State and Status   (Operator & Expert)

- ExecLowLevelCmd (Expert) : see SmarGon or DeltaTAU documentation for low level commands.

- GetFirmwareVersion   (Expert)

- GetUCodeVersion   (Expert)

- Reboot  (Expert)

- Reset   (Expert).

Figure 1 : SGonPPMACBox in Operator View

Figure 2 : SGonPPMACBox in Expert View

## ComposedAxis Management

The available operations for SGonPPMACComposedAxis Device are:

- Read current axes state,

- Change axes parameters (velocity, acceleration, deceleration, …),

- Move/stop axes (set position, stop),

- Change underlying motor state (ON/OFF),

- Parallel movements:

  - use « freeze/unfreeze » facility to set position on multiple axes at a time : the write parts of all of the axes (even those not updated) are sent to the hardware in one command.

- Start/pause/resume the CS motion program (if any defined in MotionPrograms property).

Acquisition and Control parameters:

- <label>:  axis position for the <label> axis (Operator)

- freeze (Operator)

- movingMode (Operator) : the only available value for now is 1 = LINEAR.

- velocity  (Expert)

- <label>compensationEnabled (Expert) : only available for OMEGA, CHI, PHI axes. For the other axes, the attribute is not valid.

- accelSCurve (Expert)

- acceleration  (Expert)

- deceleration  (Expert)

- <label>offset (Expert) :  offset on axis position for the <label> axis  (Expert) : only available for X, Y, Z axes. For the other axes, the attribute is not valid.

- current motion program state (Operator)

- CS status (Operator) : cf. § 3.2.

|  | Beware: Moving axes is forbidden in MOVING, FAULT or OFF state. |
| --- | --- |

Available Commands:

- Init, State and Status (Operator)

- On / Off (Operator) : power ON / OFF all axes.

- Stop (Operator)

- Start/pause/resume a motion program (Operator)

- InitializeReferencePosition (Operator) : execute referencing program. The referencing status is written in device status (cf. § 3.2)

- Off_omega (Expert) : power OFF omega axis.

- Off_others (Expert) : power OFF all axes except OMEGA.

|  | Usage tip: The user can disable one or more commands for one axis by setting the DisabledCmds property (see § 2.3.2). |
| --- | --- |

Figure 3 : SGonPPMACComposedAxis in Operator View

Figure 4 : SGonPPMACComposedAxis added attributes in Expert View

# ANNEXES

## Details of box device status

The box device status is composed of:

- Hardware connections state (open, closed, error)

- controller status read back from hardware (only the set bits are written in status).

## Details of axis device status

The composed axis device status is composed of:

- Hardware connection state (OPEN, CLOSED, ERROR…)

- CS status read back from hardware (only the set bits are written in status)

- Motors’ status (only the set bits are written in status)

- Referencing status (unknown, ok, in progress, in error)

- Calibration status (ok, ko)

- Last started motion program name.

Example:

Connection is open, Device is ready

Detailed status:

--------------------------------------

CS_STATUS :

+ In position

+ Amplifier enabled

+ Closed-loop mode

+ Home complete

+ Desired velocity zero

--------------------------------------

Motors Status :

Motor #1 status:

Home complete

Desired velocity zero

Closed-loop mode

Amplifier enabled

In position

Motor #2 status:

Home complete

Desired velocity zero

Closed-loop mode

Amplifier enabled

In position

Motor #3 status:

Home complete

Desired velocity zero

Closed-loop mode

Amplifier enabled

In position

Motor #4 status:

Home complete

Desired velocity zero

Closed-loop mode

Amplifier enabled

In position

Motor #5 status:

Home complete

Desired velocity zero

Closed-loop mode

Amplifier enabled

In position

Motor #6 status:

Home complete

Desired velocity zero

Closed-loop mode

Amplifier enabled

In position

Phase reference established

--------------------------------------

Referencing Status :

+ Referenced successfully

+ Calibration table VALID

Last started motion program: -- none --