## Modbus
| Parameter ( short name ) | ABC2.2 Modbus protocol (sw version 3.0) | [5 empty] | Questions / Comments
|  | Description | Register address |  | Number of registers | Format | Type ( Read (R) / Write (W) )
Input registers | Firmware Version | Firmware Version (for example: 'RS1.30') | 30001 | 0 | 6 | uint16 | R
| Serial number | Serial number (for example''00049') | 30007 | 6 | 2 | ulong | R
| vial type code | vial type code (0: no vial, 1: "SDD-100" ) | 30009 | 8 | 1 | uint16 | R
| vial serial # | 2x4 digits of 'vial serial # (for example:''SDD-0123-4567'-> 0123, 4567) | 30010 | 9 | 2 | uint16 | R
| calibration | = device calibration factor * vial calibration factor / 100 | 30012 | 11 | 2 | ulong | R
| initial loading | initial loading (=initial number of bubbles) | 30014 | 13 | 1 | uint16 | R
| Drops at test start | Number of drops at test start | 30015 | 14 | 1 | uint16 | R
| Start time | start time of measurement ( 7 char: year, month, day, hour, minutes, seconds, centiseconds ) | 30016 | 15 | 7 | uint16 | R
| Temperature | temperature in Celsius * 100 | 30023 | 22 | 1 | uint16 | R
| battery charge | % charge remaining (0-100, 255-> AC) | 30024 | 23 | 1 | uint16 | R
| Measurement number | Measurement number | 30025 | 24 | 1 | uint16 | R
| Total dose | Dose since the beginning of history (in nSv) | 30026 | 25 | 2 | ulong | R
| Dose since reset | Dose since reset (in nSv) | 30028 | 27 | 2 | ulong | R
| Total pops | Total number of pops (counts) | 30030 | 29 | 1 | uint16 | R
| Pops since reset | Number of pops since dose reset (count) | 30031 | 30 | 1 | uint16 | R
| Drops remaining in vial | Number of drops remaining in vial | 30032 | 31 | 1 | uint16 | R
| external alarm firing state | External alarm firing state (0: no, >=1: yes) | 30033 | 32 | 1 | uint16 | R
| Dose alarm tripped | Dose alarm tripped (0: no, >=1: yes) | 30034 | 33 | 1 | uint16 | R
| device calibration factor | range: 0 - 200, default: 100 | 30035 | 34 | 1 | uint16 | R
| Vial calibration factor | range: 0 - 255, default: 127 | 30036 | 35 | 1 | uint16 | R
| date of last device reset | last reset of device: manual or auto ( 7 char: year, month, day, hour, minutes, seconds, centiseconds ) | 30037 | 36 | 7 | uint16 | R
| date of vial reset | date of vial reset | 30044 | 43 | 5 | uint16 | R
| display version | display firmware version (eg. 1.4) | 30049 | 48 | 6 | uint16 | R
| Temperature warning state | The state of the temperature alarm (0: temperature is below the temperature alarm threshold, 1: temp is too high) | 30055 | 54 | 1 | uint16 | R | new
| Vial alarm state | 3 warning levels based on "Remaining drops/Initial loading" rate. 0 (Alarm off): >66%, 1 (warning Level 1): 20-66%, 2 (warning level 2): <20% | 30056 | 55 | 1 | uint16 | R | new
Holding registers | Dose alarm mode | Dose alarm mode (0: alarm off, 1 = alarm on, 2 = alarm on + external) | 20001 | 0 | 1 | uint16 | R + W
| Dose alarm level | Dose alarm level (in nSv) | 20002 | 1 | 2 | ulong | R + W
| --no info-- | Always = 0. Kept only as a place holder for compatibility reasons. | 20004 | 3 | 1 | uint16 | R | new
| Dose alarm displayed/muted | RABC-1: Dose alarm displayed (0: no, >=1: yes) | 20005 | 4 | 1 | uint16 | R + W
RABC-2: Dose alarm muted (0: no, >=1: yes)
| Dose reset | Dose reset (>0: dose should be reset, 0: dose just resetted) | 20006 | 5 | 1 | uint16 | R + W
| Measurement number++ | New measurement number (1: Increase measurement number by 1., 0: it is done (allowed range: 0-99, after 99 it goes back to 0)) | 20007 | 6 | 1 | uint16 | R + W
| Current time | current time (7 char: year, month, day, hour, minutes, seconds, centiseconds) | 20008 | 7 | 7 | uint16 | R + W | new
| Device calibration factor | range: 0 - 200, default: 100 | 20015 | 14 | 1 | uint16 | R + W | new
| Vial calibration factor | range: 0 - 255, default: 127 | 20016 | 15 | 1 | uint16 | R + W | new
| Temperature threshold | temperature alarm threshold in Celsius | 20017 | 16 | 1 | uint16 | R + W | new